## Data
Calibration data for the humditiy dependence of Picarro 1 (Mako) isotope ratio measurements, for all years of the ORACLES campaign. The data from fall 2016 is from David in the field; data from May and June 2017 is from Dean in the lab; fall 2017 data and fall 2018 are from Dean in the field. For isotope measurments, errors are standard deviations on the mean; for humidity the errors are standard deviations. For the fall 2018 data the original file created in the field are standard deviations, so they will not match values quoted here. They are calculated here by dividing the duration of a measurment by 30s, and using the resulting N in the formula std/N. For the 'Quality' column, G=good and OL=outlier
d18O* and dD* are generated by subtracting out the delta value measured at the highest humidity for that day. I.e. the delta value at highest humdity will be shifted to 0, and assumed to be the 'true' delta value of vapor being sampled; therefore d18O* and dD* become the deviations from 'true'.
date | d18O_StdError | d18O_permil | dD_StdError | dD_permil | h2o_ppmv | h2o_std | d18O*_permil | dD*_permil | Quality
2016-09-09 |  | -24.72 |  | -167.3 | 14370 |  | 0 | 0 | G
2016-09-09 |  | -24.5 |  | -167.9 | 5418 |  | 0.22 | -0.6 | G
2016-09-09 |  | -24.11 |  | -166.8 | 3076 |  | 0.61 | 0.5 | G
2016-09-09 |  | -23.95 |  | -166.5 | 2843 |  | 0.77 | 0.8 | G
2016-09-09 |  | -22.36 |  | -160.4 | 596 |  | 2.36 | 6.9 | G
2016-09-09 |  | -19.18 |  | -152.2 | 301.7 |  | 5.54 | 15.1 | G
2016-09-09 |  | -12.52 |  | -120.1 | 103.3 |  | 12.2 | 47.2 | G
2016-09-09 |  | -23.17 |  | -161.4 | 896.7 |  | 1.55 | 5.9 | G
2016-09-09 |  | -23.75 |  | -165.6 | 8252 |  | 0.97 | 1.7 | G
2016-09-09 |  | -24.67 |  | -165.8 | 14330 |  | 0.05 | 1.5 | G
2016-09-21 |  | -25.79 |  | -168.4 | 4033 | 140.2 |  |  | G
2016-09-21 |  | -25.9 |  | -168.5 | 8168 | 258.5 |  |  | G
2016-09-21 |  | -25.16 |  | -167.2 | 14250 | 12.2 |  |  | G
2016-09-21 |  | -25.85 |  | -167.8 | 8877 | 166.9 |  |  | G
2017-05-26 |  |  | 23.2596 | -188.6878 | 751.7735 | 45.8728 |  | -69.621 | OL
2017-05-26 |  |  | 11.2585 | -142.9875 | 1395.9773 | 52.125 |  | -23.9207 | G
2017-05-26 | 0.6358 | -20.3999 | 10.8629 | -129.1753 | 1726.1019 | 68.4054 | -2.6356 | -10.1085 | G
2017-05-26 | 0.58142 | -20.1249 | 10.039 | -128.0107 | 1876.0672 | 50.0019 | -2.3606 | -8.9439 | G
2017-05-26 | 0.37126 | -19.0729 | 6.758 | -126.8057 | 2826.4404 | 38.5586 | -1.3086 | -7.7389 | G
2017-05-26 | 0.25421 | -18.3521 | 4.7226 | -124.5974 | 3952.6997 | 51.6666 | -0.58771 | -5.5306 | G
2017-05-26 | 0.23718 | -18.147 | 4.3359 | -122.9024 | 4912.4529 | 58.7696 | -0.38266 | -3.8356 | G
2017-05-26 | 0.20256 | -17.9927 | 3.6245 | -122.067 | 5825.7448 | 54.2304 | -0.22839 | -3.0001 | G
2017-05-26 | 0.1818 | -17.8309 | 3.1194 | -121.8781 | 6925.9231 | 27.9927 | -0.066586 | -2.8113 | G
2017-05-26 | 0.16179 | -17.7643 | 2.5958 | -121.2846 | 7869.3177 | 30.2198 | 7.8152e-05 | -2.2178 | G
2017-05-26 | 0.1422 | -17.7148 | 2.3935 | -120.4044 | 8832.0164 | 22.6703 | 0.04957 | -1.3376 | G
2017-05-26 | 0.12639 | -17.7022 | 1.8646 | -119.7617 | 10870.2115 | 11.7103 | 0.062145 | -0.69482 | G
2017-05-26 | 0.11426 | -17.7644 | 1.5998 | -119.0668 | 12904.5506 | 16.0371 | 0 | 0 | G
2017-05-29 | 0.78407 | -19.9445 | 12.8466 | -134.9331 | 1278.4174 | 64.7954 | -3.6993 | -24.5409 | G
2017-05-29 | 0.43672 | -18.4577 | 7.4203 | -125.513 | 2000.2884 | 68.5947 | -2.2125 | -15.1208 | G
2017-05-29 | 0.50126 | -18.0493 | 8.7827 | -123.508 | 2322.0299 | 74.5577 | -1.8041 | -13.1159 | G
2017-05-29 | 0.42459 | -17.7313 | 7.3177 | -122.2288 | 2489.3062 | 62.4678 | -1.4861 | -11.8366 | G
2017-05-29 | 0.29346 | -17.1414 | 5.3392 | -119.1556 | 3561.7475 | 39.9082 | -0.89622 | -8.7634 | G
2017-05-29 | 0.21312 | -16.7519 | 3.9771 | -117.2591 | 4773.8541 | 57.1041 | -0.5067 | -6.8669 | G
2017-05-29 | 0.20505 | -16.5673 | 3.9681 | -116.0728 | 5802.917 | 47.4024 | -0.3221 | -5.6806 | G
2017-05-29 | 0.17278 | -16.5192 | 3.2096 | -114.7587 | 6829.4742 | 50.1602 | -0.27401 | -4.3665 | G
2017-05-29 | 0.1598 | -16.4091 | 2.8417 | -114.2579 | 8033.0114 | 23.2907 | -0.16395 | -3.8657 | G
2017-05-29 | 0.14306 | -16.319 | 2.3986 | -113.5019 | 9064.3436 | 17.8592 | -0.073834 | -3.1097 | G
2017-05-29 | 0.14082 | -16.2618 | 2.1885 | -113.3506 | 10092.3733 | 20.2878 | -0.016612 | -2.9584 | G
2017-05-29 | 0.12181 | -16.2626 | 1.8245 | -112.583 | 12315.0785 | 13.2488 | -0.017422 | -2.1908 | G
2017-05-29 | 0.11026 | -16.1975 | 1.4614 | -111.3857 | 14566.9401 | 19.0423 | 0.047722 | -0.99356 | G
2017-05-29 | 0.10237 | -16.3123 | 1.3589 | -111.0791 | 16789.2612 | 15.1352 | -0.067096 | -0.68695 | G
2017-05-29 | 0.077719 | -16.2452 | 0.85196 | -110.3922 | 19060.5042 | 11.6497 | 0 | 0 | G
2017-05-29 | 0.1264 | -16.2351 | 1.8943 | -111.2897 | 12334.5461 | 16.746 | 0.01008 | -0.89748 | G
2017-05-29 | 0.23517 | -16.5716 | 4.2756 | -113.1585 | 4748.6877 | 55.1886 | -0.32639 | -2.7663 | G
2017-05-29 | 0.72121 | -19.2923 | 12.2644 | -122.9208 | 1273.0874 | 51.7747 | -3.0471 | -12.5286 | G
2017-06-01 | 3.1825 | -26.0148 | 52.7874 | -155.536 | 436.1682 | 28.8637 | -11.2211 | -57.3513 | G
2017-06-01 | 1.8753 | -21.5827 | 29.3579 | -130.8389 | 731.4851 | 29.936 | -6.789 | -32.6542 | G
2017-06-01 | 1.3378 | -19.43 | 21.3827 | -122.3515 | 1022.8527 | 43.4008 | -4.6363 | -24.1668 | G
2017-06-01 | 0.79732 | -17.4985 | 14.4134 | -116.0373 | 1657.6303 | 65.4443 | -2.7048 | -17.8526 | G
2017-06-01 | 0.69009 | -17.1671 | 12.3649 | -112.6 | 1819.4215 | 56.3339 | -2.3734 | -14.4153 | G
2017-06-01 | 0.53293 | -16.493 | 10.0125 | -110.7281 | 2310.4824 | 38.2421 | -1.6993 | -12.5434 | G
2017-06-01 | 0.37297 | -15.986 | 6.8757 | -105.7836 | 3106.1249 | 51.2463 | -1.1923 | -7.5989 | G
2017-06-01 | 0.31572 | -15.6898 | 5.7603 | -104.6256 | 3866.0673 | 50.5775 | -0.89609 | -6.4409 | G
2017-06-01 | 0.2592 | -15.5719 | 4.8076 | -101.8304 | 4676.9447 | 25.6451 | -0.77823 | -3.6457 | G
2017-06-01 | 0.23722 | -15.2427 | 4.4898 | -101.6678 | 5951.4908 | 26.8388 | -0.44897 | -3.4831 | G
2017-06-01 | 0.1912 | -15.0099 | 3.3465 | -100.9839 | 8019.0295 | 20.3517 | -0.21615 | -2.7992 | G
2017-06-01 | 0.15065 | -14.8509 | 2.6516 | -100.2589 | 10076.3012 | 23.1191 | -0.057162 | -2.0742 | G
2017-06-01 | 0.10699 | -14.821 | 1.6151 | -99.2998 | 13259.0651 | 16.6364 | -0.0273 | -1.1151 | G
2017-06-01 | 0.10135 | -14.753 | 1.3421 | -98.5877 | 15329.0871 | 14.9044 | 0.040724 | -0.40299 | G
2017-06-01 | 0.10678 | -14.7937 | 1.2899 | -98.1847 | 17408.4285 | 13.3212 | 0 | 0 | G
2017-06-01 | 0.10695 | -14.6953 | 1.4624 | -98.0283 | 15341.4673 | 13.9279 | 0.098394 | 0.1564 | G
2017-06-01 | 0.10946 | -14.7073 | 1.5399 | -98.4852 | 13279.9323 | 17.5023 | 0.08637 | -0.30047 | G
2017-06-01 | 0.16029 | -14.6243 | 2.7425 | -99.242 | 10114.6485 | 13.9606 | 0.16943 | -1.0573 | G
2017-06-01 | 0.19109 | -14.6711 | 3.2567 | -100.3305 | 8048.3291 | 22.672 | 0.1226 | -2.1458 | G
2017-06-01 | 0.23717 | -14.7131 | 4.3867 | -101.3842 | 6010.0879 | 27.2385 | 0.080582 | -3.1995 | G
2017-06-01 | 0.24902 | -15.0027 | 4.6982 | -101.7656 | 4719.6808 | 32.514 | -0.20898 | -3.5809 | G
2017-06-01 | 0.2821 | -15.3323 | 5.4788 | -102.0956 | 3923.989 | 56.4174 | -0.5386 | -3.9109 | G
2017-06-01 | 0.3688 | -15.5286 | 6.8111 | -105.0878 | 3135.3697 | 56.1046 | -0.73486 | -6.9031 | G
2017-06-01 | 0.52294 | -16.1668 | 9.6485 | -105.9225 | 2324.0529 | 44.0772 | -1.3731 | -7.7378 | G
2017-06-01 | 0.67282 | -16.8796 | 12.0023 | -109.3213 | 1832.9409 | 46.9407 | -2.0859 | -11.1366 | G
2017-08-16 |  | -18.8236 |  | -129.9396 | 17635.2173 |  | 0.029186 | -1.247 | G
2017-08-16 |  | -18.8129 |  | -129.9672 | 14945.2302 |  | 0.039863 | -1.2746 | G
2017-08-16 |  | -18.863 |  | -130.4115 | 12848.0463 |  | -0.010246 | -1.7189 | G
2017-08-16 |  | -18.7431 |  | -130.1688 | 10082.0278 |  | 0.10959 | -1.4762 | G
2017-08-16 |  | -18.8001 |  | -131.0585 | 7231.3144 |  | 0.052647 | -2.3659 | G
2017-08-16 |  | -18.8442 |  | -130.8622 | 5523.2741 |  | 0.0084917 | -2.1696 | G
2017-08-16 |  | -19.1249 |  | -131.9402 | 4102.3261 |  | -0.27219 | -3.2476 | G
2017-08-16 |  | -19.4279 |  | -132.2125 | 2708.8488 |  | -0.57513 | -3.5199 | G
2017-08-16 |  | -19.7316 |  | -134.4764 | 1994.6597 |  | -0.87882 | -5.7838 | G
2017-08-16 |  | -19.3712 |  | -133.5994 | 2713.5459 |  | -0.51849 | -4.9068 | G
2017-08-16 |  | -19.1922 |  | -131.6403 | 4120.8042 |  | -0.33945 | -2.9477 | G
2017-08-16 |  | -18.9159 |  | -130.7862 | 5490.5253 |  | -0.063186 | -2.0936 | G
2017-08-16 |  | -18.8664 |  | -130.7396 | 7220.8158 |  | -0.013677 | -2.047 | G
2017-08-16 |  | -18.7561 |  | -130.3546 | 10092.0078 |  | 0.096642 | -1.662 | G
2017-08-16 |  | -18.8322 |  | -129.356 | 12834.2532 |  | 0.020523 | -0.66339 | G
2017-08-16 |  | -18.7812 |  | -129.4309 | 14958.9504 |  | 0.071539 | -0.7383 | G
2017-08-16 |  | -18.8527 |  | -128.6926 | 17676.814 |  | 0 | 0 | G
2017-08-16 |  | -18.8828 |  | -128.8086 | 14953.5925 |  | -0.030045 | -0.11603 | G
2017-08-16 |  | -18.7765 |  | -129.1394 | 12846.0174 |  | 0.076225 | -0.44676 | G
2017-08-16 |  | -18.7308 |  | -129.4023 | 10102.6538 |  | 0.12191 | -0.70974 | G
2017-08-16 |  | -18.7524 |  | -129.9566 | 7266.064 |  | 0.10035 | -1.264 | G
2017-08-16 |  | -18.8289 |  | -130.1858 | 5498.6063 |  | 0.023884 | -1.4931 | G
2017-08-16 |  | -18.9503 |  | -131.6424 | 4102.1502 |  | -0.097516 | -2.9498 | G
2017-08-16 |  | -19.1899 |  | -132.8351 | 2711.68 |  | -0.33716 | -4.1425 | G
2017-08-27 | 0.14396 | -23.7976 | 1.7744 | -168.3061 | 17605.1912 | 19.4507 | 0 | 0 | G
2017-08-27 | 0.16586 | -23.906 | 2.3145 | -168.3619 | 14870.7817 | 21.8873 | -0.10836 | -0.055812 | G
2017-08-27 | 0.16628 | -23.9025 | 2.512 | -168.9949 | 12688.5606 | 20.0044 | -0.10487 | -0.68879 | G
2017-08-27 | 0.18017 | -23.8626 | 3.418 | -169.4321 | 9862.4881 | 51.5855 | -0.064938 | -1.126 | G
2017-08-27 | 0.22496 | -23.9341 | 4.7504 | -170.0029 | 6899.8681 | 45.8627 | -0.13649 | -1.6968 | G
2017-08-27 | 0.30965 | -23.9407 | 6.1553 | -171.3636 | 5124.4917 | 44.169 | -0.14307 | -3.0575 | G
2017-08-27 | 0.38197 | -24.0847 | 7.8929 | -172.6543 | 3712.964 | 30.6795 | -0.28709 | -4.3482 | G
2017-08-27 | 0.39717 | -24.2482 | 8.4143 | -171.9723 | 3151.8483 | 37.6768 | -0.4506 | -3.6662 | G
2017-08-27 | 0.47315 | -24.5917 | 9.5999 | -172.1721 | 2425.811 | 45.2966 | -0.79407 | -3.866 | G
2017-08-27 | 0.68286 | -25.1774 | 13.5297 | -172.1806 | 1761.6791 | 36.8341 | -1.3797 | -3.8746 | G
2017-08-27 | 0.61729 | -24.772 | 12.2083 | -172.8476 | 2386.7408 | 38.8516 | -0.97435 | -4.5416 | G
2017-08-27 | 0.39971 | -24.2924 | 8.2225 | -171.7789 | 3123.4948 | 46.5262 | -0.49482 | -3.4728 | G
2017-08-27 | 0.28044 | -24.0855 | 5.6484 | -171.4145 | 4439.8459 | 51.6284 | -0.28788 | -3.1084 | G
2017-08-27 | 0.2601 | -23.8855 | 5.4238 | -171.1479 | 5785.8301 | 32.6821 | -0.087883 | -2.8418 | G
2017-08-27 | 0.21541 | -23.9159 | 3.8694 | -170.0839 | 7435.764 | 39.4789 | -0.11825 | -1.7778 | G
2017-08-27 | 0.19144 | -23.87 | 3.391 | -170.1247 | 10222.579 | 27.4298 | -0.072433 | -1.8187 | G
2017-08-27 | 0.16064 | -23.9575 | 2.4575 | -169.554 | 12887.5579 | 14.6121 | -0.15988 | -1.2479 | G
2017-08-27 | 0.16364 | -23.8347 | 2.169 | -169.2971 | 14953.1694 | 11.4954 | -0.037108 | -0.99102 | G
2018-10-19 | 0.56 | -22.3 | 10.5 | -161 | 600 | 60 | -5.3 | -29 | G
2018-10-19 | 0.45 | -20.8 | 8.2 | -148 | 890 | 62 | -3.8 | -16 | G
2018-10-19 | 0.2 | -18.5 | 3.6 | -140 | 2000 | 88 | -1.5 | -8 | G
2018-10-19 | 0.09 | -17.5 | 1.7 | -135.4 | 4300 | 100 | -0.5 | -3.4 | G
2018-10-19 | 0.1 | -17.2 | 2 | -135.5 | 6000 | 95 | -0.19999 | -3.5 | G
2018-10-19 | 0.08 | -17.2 | 1.3 | -134 | 8100 | 70 | -0.199 | -2 | G
2018-10-19 | 0.07 | -16.9 | 1 | -133.4 | 11700 | 30 | 0.1 | -1.4 | G
2018-10-19 | 0.1 | -16.9 | 1 | -132.5 | 15000 | 30 | 0.1 | -0.5 | G
2018-10-19 | 0.1 | -17 | 1 | -132 | 17600 | 25 | 0 | 0 | G

## Figures
Estimate of uncertainty in humidity dependence calibration from spread in calibration data over all years (excluding Fall 2016)
Humidity, ppmv | +/- dD error, permil | +/- d18O error, permil
20000 | 1 | 0.15
15000 | 1 | 0.2
10000 | 1.5 | 0.2
5000 | 2.5 | 0.35
2500 | 4.5 | 0.5
1000 | 7 | 1